FYS 4250

20 questions to each lecture

Lecture 1

- Can you make a simple sketch of a simple ear-thermometer and explain the principles of function?

- Can you make a simple sketch of a kidney stone crusher and explain the principles of function?

- What are the main risks for the patient when using a kidney stone crusher, and how can these be minimized?

- What is a cochlea implant and how does it work?

- How is the human ear able to discriminate between different frequencies (Frequency tuning) and how is the cochlea implant able to do the same discrimination?

- Understanding speech is the ultimate goal and the younger the patient, the more effective the cochlea implant will be. Why?

- What is a cochlea implant and how does it work?

- What requirements are typical for a cochlea implant amplifier

- Will a cochlea implant be effective if the patient has an injury in the cochlear nerve?

- What is a demand pacemaker and how is it designed?

- When a patient has a problem with the AV-node and intermittent discontinuation in the sinus node?

- What is the advantages/disadvantages with a unipolar vs a bipolar pacemaker?

- Most modern pacemakers are bipolar, why?

- What are the two main types of electrodes, and what is the difference between them?

- Why is polarization a topic when it comes to pacemaker electrodes?

- What is a half-cell potential and what are the theoretical types of electrodes?

- Describe the basic components of the half-cell overpotential

- When the frequency of a measured signal is increasing, how will that influence the half-cell potential?

- Why cannot regular hemodialysis be used to clean the blood from toxins during a liver failure?

Lecture 2

- Can you describe the circulatory system, the main parts of it and the function of the most important parts?

- How are the different parts of the heart synchronized?

- If we measure an elevated pressure in the right atrium, what could possibly be wrong and which devices could we use to fix the problem?

- What is an action potential and why is the action potential limited in time?

- What is the main purpose of the AV-node, and what happens if the SA-node is not working as expected?

- How is it possible to diagnose the heart by measuring electricity and what kind of equipment do you need?

- Why is the ECG-repolarization positive?

- What is an Einthoven triangle and how is it used?

- What is carrying the current in the body, and which problems can occur when we will transfer this current to cables outside the body?

- What is the voltage and frequency-range for ECG-signals?

- When the frequency-range of an ECG-device is in the range of 0-100 Hz, what would the ECG-signal look like on the monitor?

- What is the Wilson-terminal, what is the advantage and what is the disadvantage with the method?

- What are augmented leads, what is the advantage and how is the lead vector different from the regular Einthoven triangle?

- What is CMRR in an ECG-amplifier and why is it important?

- How can you minimize the noise from magnetic field pickup when acquiring the ECG-signals?

- Make a simple sketch of an ECG-amplifier and discuss briefly the main parts of it

- Describe a method for actively suppressing electromagnetic noise in an ECG-setup

- Is it a built-in electrical safety system in driven right leg?

- How many leads do you need for a 12-lead ECG, where are they located and which electrodes belong to which leads?

- Why is it almost impossible to design an ECG-monitor without using operational amplifiers?

- Why is it important that the ECG-system does not "load" the patient (= there is no current from the patient through the electrodes)

Lecture 3

- Describe what blood pressures you would expect from an arterial blood pressure monitor, and which you would expect inside the different parts of the heart

- What is the difference between invasive blood pressure measurements and non-invasive measurements?

- Describe the most common non-invasive methods for blood pressure measurements and their main advantages/disadvantages and sources of error

- Which of the invasive methods are able to measure absolute pressure, and which can measure relative pressure?

- How will an air bubble affect the blood pressure measurement from an invasive blood pressure measurement system?

- How can a wheatstone-bridge be useful for blood pressure measurements?

- What are the main advantages and disadvantages with intra-vascular blood pressure measurements?

- How can you measure the system characteristics of an invasive blood pressure system?

- What are the main sources of error for a wrist blood pressure measurement system, and what are the advantages?

- What is a stethoscope, how does it work and what is a typical frequencyresponse?

- How can you measure the blood pressure by means of a stethoscope, a sphygmomanometer and a cuff?

- Why is an automatic bloodpressure monitor with a sphygmomanometer, cuff and a microphone a bad idea?

- Why is it impossible to measure the diastolic pressure with an oscillometric method?

- What are the typical error sources for sphygmomanometers?

- EXPERT: Which non-invasive blood pressure measurement method would you recommend for measuring the blood pressure of a patient suffering from muscle-atrophy?

- What kind of device would you use for diagnosing glaucoma?

- How do you know that you have applied the right pressure to a manual tonometer?

- What are the basic principles of an automatic contact-tonometer?

- Which tonometry-method has the least risk of patient injury?

- What are the basic principles of an automatic non-contact tonometer?

Lecture 4

- What is the oxyhemoglobin dissociation curve, and what is the implication for the patient?

- What is saturation, what is the unit and how can we calculate it?

- Describe the cycle of breathing, what happens?

- If a person inhales more than the regular resting volume of 0,5 liter of air in the same length of time, would intra-alveolar pressure increase or decrease (become more negative)?

- Make a sketch of a pH electrode, describe the different parts and what requirements would you need from an amplifier?

- What is a fiberoptic pH-sensor, and what are the major advantages/disadvantages/sources of error for this?

- Why do we use green and red light for optical absorbance using pH sensitive dyes?

- What is a fluorescent dyesensor, and what do we mean with "irreversible" in this context?

- How do we read out the pH from a fluorescent dyesensor?

- Describe the main parts and function of a pCO2 electrode, and mention the most important sources of error

- Why is the core of the pCO2-electrode made of a pH-electrode?

- Describe the pO2 electrode, the main parts and the function and sources of error

- Why do we need a voltage source in the pO2 electrode, and why is it 0,7 Volt?

- Describe a simple system for measuring saturation, and explain how pulse oximetry is made

- How can we be sure that a pulse oximeter does not measure saturation in venous blood?

- If you would like to measure the blood flow with a thermistor, why is it a bad idea to have constant current through the resistor?

- Show a simple circuit that can provide constant temperature in a thermistor measuring the blood flow, and explain how you compensate for temperature variations in the blood.

- How can you diagnose a venous thrombosis (a clot in the vene) in a leg without imaging- or invasive methods?

- What are the main assumptions that is the basis for impedance plethysmography

- What are the main problems with a two-electrode measurement?

Lecture 5

- How can we control the oxygen-level when a patient is ventilated with a self-inflating O2-concentrations?

- How can we avoid wasting costly anesthetics when ventilating patients with mechanical ventilation, what are the main risks with the method and how can we minimize this risk?

- What is microdialysis and how can this method be used for finding concentrations in the extracellular fluid in the brain?

- What is the difference between low-flow rate and no-net flux?

- Why is always a volume of air trapped in the lungs?

- What are the main parts of a servocontrolled ventilator, and what would happen if we removed the expiratory valve?

- What is volume-controlled ventilation, and how is the servo-ventilator acting to provide the volume-controlled mode?

- What is pressure-controlled ventilation, and what factors affect the volume in this mode?

- If you have a patient where you suspect high airway resistance, which mode would you select?

- What are the main advantages with each mode, and what are the main disadvantages?

- What happens if the inspiration time Ti is too short during inspiration?

- What happens if the expiration time Te is too short during expiration?

- What is peep, and how can the servo-ventilator achieve a decent peep?

- How will low compliance influence on mechanical ventilation, and what would the ventilator-operator do to minimize the problem if the compliance decreases?

- What are the main advantages and disadvantages with each ventilator mode?

- What is a sidestream measurement, and what is the main advantages and disadvantages with this method

- What is a typical error with a spectrophotometric method, and how can we minimize the problem?

- What are the main sources of error when using a paramagnetic oxygen analyzer?

- What is the main advantage with a full-body plethysmograph?

- Explain the principles of mechanical hemodialysis, and explain why there is a pump generating an under-pressure in the tube that leads the waste to the sink

Lecture 6

- Describe how the X-ray tube is designed

- What is the collimator light good for?

- Why is the mass attenuation coefficient important?

- What is the difference between “Brems-strahlung” and characteristic radiation?

- Why do we need a filter before the photons are leaving the x-ray tube?

- What Is secondary radiation, what is the problem with it and how can we avoid it?

- What kind of focal spot would you prefer if you were taking an X-ray image of a very small, bony structure in the wrist?

- What are the main sources of error, and how can we minimize each?

- What is a modulation transfer function and how can we derive it for an x-ray system?

- What are the three main ways of protecting against ionizing radiation?

- What is a contrast agent and what is the purpose of it?

- What are the basic principles of a CT-scanner?

- What is a simple back-projection and what is the main problem with the method?

- What kind of filter is used to minimize the filtered backprojection error?

- What is an iterative projection?

- Explain the basic stages of an iterative projection

- What is special with a CT-tube?

- What defines the resolution of a CT-image?

- How can a proper CT-image of the heart be taken without having an excessive dose?

- What kind of energy is used for mammography X-rays?

Lecture 7

- What is the principle of fluorescence spectroscopy?

- Describe the basic function of a photomultiplier and explain why it is a negative voltage at the photocatode

- Why is electrosurgery relatively safe for the patient, even if quite high currents are sent into the patient?

- Explain the basic setup of an electrosurgery unit and describe the safety measures used to maximize patient safety

- What is the difference between monopolar and bipolar electrosurgery, and which one is the safest for the patien?

- What is the best method for avoiding or minimizing the antenna effect (= current Is jumping from the electrosurgery cables/electrode to nearby metal objects)?

- How can you take control with the current path?

- What is the law of physics that the electromagnetic flowmeter is based on, and why is it a correct assumption to say that e = BLu?

- How is the sensitivity field for the electrodes shaped, and what are the main error sources for the electromagnetic flowmeter?

- What are the practical limitations for the electromagnetic flowmeter?

- What is the best, a DC flowmeter or an AC flowmeter, and why?

- What is the main problem with an AC flowmeter and what are the two methods for counteracting this problem?

- What is the basic function of a quadrature suppression 	flowmeter?

- What happens if the quadrature generator is 180 degrees phase shifted instead of 90?

- What are the main parts of the blood, and how can we separate it?

- What is a cell-saver, why do we use it, and why do we remove almost everything except water and red blood cells?

- What tools do you need for a laparascopic procedure?

- Why is it important that you can switch a laparascopic procedure to open surgery?

- Why is a laparascopic procedure better for the patient?

- What is the main risk-factors for a laparascopic procedure?

Lecture 8

- What is a pulmonary artery catheter (PAC), and what is this used for?

- How can we place the PAC in the right place of the circulatory system?

- What is the most acute risk of the PAC?

- What is Fick’s method, what are the main advantages and why is it obsolete?

- What is an indicator dilution method, and how can we use this?

- What is the main reason for using thermodilution instead of dye-dilution, making thermodilution the preferred indicator method?

- Why is it important that indocyanine green has its optical absorption peak at 805 nm?

- What are the properties of a piezoelectric crystal that makes it suitable for ultrasound measures?

- Describe the design and function of an ultrasound doppler flowmeter

- What are the main sources of error?

- The amplifier will use both a low pass filter and a high pass filter, explain why?

- Why do we need a quadrature-phase detector?

- Make a simple sketch of the quadrature detector and explain each element in the circuit

- Compare the electromagnetic flowmeter and the ultrasound doppler flowmeter. What are the main advantages and disadvantages with each?

- Why is a cerebral oximeter measuring a mix of venous and arterial blood, instead of arterial blood only, as the regular pulse oximeter is doing?

- Lambert-Beers law describes the physical properties of absorption of light, indicating a source on one side of the object and a detector on the other side. However, a cerebral oximeter has both detectors and light sources on the same side. How can we still use the equation?

- What is the basic mechanism of ultrasound surgery?

- In which cases would we prefer ultrasound surgery before regular electrosurgery?

- How can a surgical tracking system make sure that every tool used is correctly aligned in the image shown to the surgeon?

- What are the main risks of intraoperative navigation?

Lecture 9

- What is the traditional way of measuring stages of anesthesia, and how can they separate them?

- Make a simple sketch of an anesthesia machine and describe how it works

- What are the main functions of the anesthesia workstations and what are the main concerns?

- What is the function of a vaporizer on the anesthesia workstation, what is the main problem with it and how can this be compensated?

- What are the main concerns with the scavenging system, and what is special about the valve on this circuit?

- What could happen if the positive pressure relief valve (APL) from the scavenging system does not work properly?

- What could happen if the negative pressure relief valve from the scavenging system does not work properly?

- Why is the APL valve used in the manual mode only?

- What is the main function of the CO2-absorber (soda lime)?

- What is the function of the HME-unit, and what kind of problem could this unit introduce?

- There are two important flowmeters on the anesthesia machine, where are they, and what is the function?

- How it possible to measure anesthesia automatically, and what are are the main sources of error?

- What is a slave regulator, and what is the difference between a two room and a three room slave regulator?

- What can happen if the O2-flush is used for a long time? (It is type-dependent)

- If a patient connected to a BIS-monitor develop fever during surgery, would you expect the BIS-monitor to show a higher BIS-Score or a lower BIS-Score than the real value, or does it calculate the real value anyway?

- EXPERT: Except from BIS-Score, which monitoring parameters could reveal an ongoing awareness?

- EXPERT: Giving anesthesia is to put patients into chemical comas, but there are reports of   30 - 40 000 yearly cases of awareness (the patient wakes up) during anesthesia in the US. In case of awareness, will the patient feel pain when waking up, why/why not?

- EXPERT: What happens with the anesthesia machine if the O2-sensor does not measure correctly?

- EXPERT: What is the difference between a ventilator and the ventilator part of the anesthesia workstation?

- EXPERT: During the first wave of the Covid-19 pandemic, hospitals discussed to use anesthesia workstations as ventilators due to the lack of available ventilators. Why is this problematic, and is it possible?

Lecture 10

- What is a spectrophotometer and how is it designed? (Make a sketch)

- What is an instrument for flame emission, and what is the difference compared to flame absorption? (Make a sketch)

- What is the difference between flame photometers and spectrophotometers?

- What is fluorometry and what is the advantages/disadvantages compared to a spectrophotometer?

- Give a brief overview of the gas/liquid cromatograph (GLC), what is the main advantages/disadvantages with this method?

- What is electrophoresis and what are the main advantages/disadvantages?

- What is the principle of a coulter-counter and what is it used for?

- How is the genetic information stored in a cell, and how can we take advantage of this in a polymerase chain reaction?

- What are the main three cycles of a PCR?

- What is the problem with traditional PCR?

- What is the advantage with qPCR or realtime PCR compared to traditional PCR, and how is such a device designed?

- Why is it important to know the number of copies made of the DNA?

- What is a mass spectrometer and how is it designed? (Make a simple sketch)

- Why do we often combine e.g. gas chromatography with mass spectrometry?

- How is the arriving ions in a mass spectrometer converted to an electrical signal, amplified and measured?

- What are the main advantages and disadvantages with mass spectrometry?

- What is a flow cytometer, and how is it designed?

- EXPERT: Why is Covid-19 antibodies measured in a flow cytometer, while the Covid-19 test is performed in a PCR?

- EXPERT: We do often combine two mass spectrometers, why?

- EXPERT: Why is HB converted to cyanmethemoglobin in the coulter counter?

Lecture 11

- How can we see typical signs of a myocardial infarction in an ECG-diagram?

- How many electrodes do you need for a 12-lead ECG?

- Explain the differences between the 12 leads

- What is a Percutaneous Coronary Intervention (PCI) procedure?

- Describe briefly the principle of an IABP (Intra arterial balloon pump)

- What are the main risks of an IABP and what can be used to reduce the risks?

- Which physical law describes the blood flow through a blood vessel, and how will a reduction of 50% in the diameter of a blood vessel influence on the flow?

- What is the main risks of using anti coagulants in patients?

- How can we measure the degree of coagulation, describe briefly the three most common methods?

- What is a roller pump and what are the main advantages/disadvantages?

- What is a centrifugal pump, and what are the main advantages/disadvantages?

- What is a “priming volume”?

- Would you use a roller pump or a centrifugal pump for a patient with a 24-hour need for assisted circulation?

- Why is it easy to calculate the flow with a roller pump, as opposed to a centrifugal pump?

- What is an oxygenator and what are the main parts?

- For a Heart-Lung Machine (HLM), is the venous reservoir open to air or a closed system, and why?

- What are the main risks with the HLM?

- If you should give cardiac support to a patient for some days, would you use Veno-arterial ECMO of Veno-venous ECMO

- How is it possible to make a medical suction based on over-pressure from the wall inlet?

- EXPERT: What would happen if the left ventricular vent tube were incorrectly mounted?

Lecture 12

- Describe briefly the main parts of a neuron, how can a neuron stimulate other neurons?

- Which cells in the brain contribute to the EEG-signals?

- What is the difference between temporal and spatial summation of neurons?

- Why is it important with myelin?

- What is the main problem with scalp electrodes?

- How can you make a tangential EEG measurement and how can you measure a radial EEG measurement

- What is the most important requirement for an EEG-amplifier?

- Why is the Nyquist theorem important for EEG-measurements?

- What is the Nernst equation actually describing?

- What does it mean that a cell is in a volume conductor and what is the problem with that?

- Which nerve fibers will give the largest signal?

- What is a reflex, and why is it important?

- How can you measure neural velocity?

- When measuring the neural velocity in the median nerve (in the hand), how would you set up the equipment and why is the signal at the elbow so different?

- What is an H-reflex, and how can we measure it?

- What is EMG and what is it used for?

- What is MEG, and how does it work?

- How can you measure a MEG-signal?

- Why is the fMRI preferred at the expense of MEG?

- What is ERG, how can it be measured and what can you get out of these measurements?

Lecture 13

- What is ultrasound?

- How can ultrasound be used for medical imaging?

- What is the relation between ultrasound frequency and attenuation, and what determines the frequency in an ultrasound system?

- What is a piezo crystal and how is it used in an ultrasound system?

- How can we form the ultrasound beam in order to focus at a certain distance?

- What is a phased array transducer and what is dynamic receive focusing?

- What is in order to maximize axial resolution?

- What is important in order to discriminate two objects perpendicular to the beam?

- What determines the axial, elevational and lateral resolution?

- What is the relation between ultrasound transducer frequency vs resolution when it comes to penetration depth and resolution?

- Describe briefly the most common ultrasound modi and give examples of the different applications for each of them

- What is time-gain compensation and why is it important?

- What are the main sources of error in a regular ultrasound system?

- What is Doppler ultrasound and what is it used for?

- What is the problem with Doppler ultrasound aliasing?

- What are the main advantages/disadvantages with ultrasound?

- What is focused ultrasound and what is it used for?

- What are the main advantages/disadvantages with focused ultrasound compared to traditional surgery?

- EXPERT: What frequency will compress and decompress the medium longest, 2 MHz or 5 MHz?

- What is spatial resolution and which factor will not influence on spatial resolution?

  - Frequency

  - Pulse interval

  - Time-gain compensation

  - Transducer thickness

Lecture 14

- What is the frequency ranges for the electromagnetic radiation energy from an MRI-imaging system, compared to an X-ray system?

- What are the basic physic properties we take advantage of when making an MRI-image?

- What is the Larmour-frequency and why is it important in MRI-imaging?

- What is T1 relaxation and explain why you need several different delaytimes between the two 90 degree pulses in order to find T1?

- What is T2 relaxation and what is spin-echo contrast weighting?

- EXPERT: What is the difference between T2 and T2*, and why is T2 always longer than T2*

- How can we find the T2-spin echo decay?

- Why is it important with a gradient magnetic field, and how can we apply such a field?

- How is the stationary magnet designed in order to produce the gradient magnetic fields?

- What is a frequency encode gradient (FEG) and how can we use that in order to obtain spatial image information?

- What is the task of a shimming coil?

- What is functional MRI and how is it possible to use the MRI-scanner for the functional image?

- Describe a couple of MRI-artefacts and explain why they appear

- What kind of method is used in the MRI-scanner for picture reconstruction?

- What are the main advantages/disadvantages with MRI?

- What are the main concerns when it comes to safety issues related to the MRI?

- What is a slice select gradient (SSG) and how do we use it?

- EXPERT: If the total signal of pixel A and B with the same encoded frequency is written:      S0(t) = (A+B) sin wt, does it mean that phase-encoding gradient has been applied or not?

- EXPERT: What is better with a 3T MRI scanner compared to a 1T scanner, and what is better with the 1T scanner when it comes to image quality?

- EXPERT: What happens if two different gradients are applied simultaneously?

Lecture 15

- What is the purpose of a defibrillator and make a simple sketch of a basic defibrillator

- What is important for a successful defibrillation?

- In which situations can it be important to synchronize defibrillations?

- What is a coronary angiography and what is it used for?

- Most imaging methods show anatomy which then has to be interpreted in a medical context. Which method is able to give an image of pathology (disease) directly?

- Make a simple sketch of a scintillation detector and explain how this is used in a gamma camera, and what a gamma camera is used for

- Why is it important to have a collimator in a gamma camera?

- What is a SPECT-camera and what is it used for?

- What are the principal differences between a PET-camera and a SPECT-camera?

- Give a brief explanation of a PET scanner and explain what a coincidence detection is

- Why don't we use collimators in a PET-scanner?

- How is the 511-keV converted (and amplified) to an electrical signal in the PET-scanner?

- What are the main advantages and disadvantages with 3D PET-scanner vs a 2D scanner?

- What is attenuation correction, why is it important, what is the disadvantages and how is it done in a modern scanner vs an old scanner?

- EXPERT: Is attenuation most important in a SPECT scanner or a PET scanner?

- If you want the highest resolution, should you select a SPECT or a PET-scanner?

- How is the image constructed in a SPECT/PET scanner?

- What is the major advantages/disadvantages with a PET-MRI scanner vs a PET-CT scanner?

- What is the major advantages/disadvantages with proton therapy compared to regular radiation therapy, and what is a "Bragg peak"?

- What makes the destructive effect inside the cancerous cells from a proton-beam?

Lecture 16

- What are the main parts and the main principle of a linear accelerator (Linac) used for radiation therapy?

- Make a simple sketch of a linac used for radiation therapy

- How is it possible to steer the radiation beam in a linac used for radiation therapy?

- Why do we need a bending magnet in a linac?

- Why is an X-ray unsuitable for treatment of a brain tumor?

- What are the basic principles of an infusion pump?

- What is the basic princples of a syringe pump?

- Why do we need a drop counter for an infusion pump but not a syringepump?

- What is the main challenge with a multipump connection, and why do we use it?

- Is it dangerous or not to touch a patient with one hand during defibrillation?

- What is the single fault concept?

- What is a let-go current?

- Why will the risk of releasing a ventricular fibrillation as a result of electrical shock reach the maximum after approximately a second?

- What are the main advantages/disadvantages with infusion pumps vs syringe pumps?

- Which legal regulations are regulating medical devices?

- What is the legal therm "patient environment" and what is special about this zone?

- What is the difference between macro and micro-shock, and what are the fibrillation thresholds for each situation?

- How is it possible to design an electrical medical device in such a way that the risk of                          current leakage from the amplifier is reduced almost to zero?

- What is a leakage current, and why is it dangerous?

- Describe one test that measures leakage currents, and explain which risk this test will deal with?